DTM-150

DIGITAL HALL EFFECT TESLAMETER

The DTM-150 Digital Teslameter offers accurate, high resolution measurement of magnetic flux densities.

Features:

Resolution of 1 part in 600,000

4 ranges

Maximum field 3 Tesla

7 digit display

either RS232 or IEEE488/GPIB communications.

Accuracy is 0.03% of reading + 0.006% of full scale when used with MPT-132 probe.

Because the probe characterisation information is stored in the probe assembly itself, any Group3 probe can be used with any Group3 DTM.

Two buttons on the front panel are used to select the operation of the teslameter. The following modes can be selected for the display - magnetic field, peak hold of magnetic field, AC component of field.

Digital filtering (time averaging) can be enabled to suppress short term fluctuations.

Several teslameters with the serial communications option (up to 31) can be formed into a local communications loop, all of them talked to by the one serial port on the control computer. The system can be operated in a trigger mode, where the timing of the measurements by several teslameters can be synchronised

Internal switches select serial data format, baud rate, device address, string terminators, filtering, gauss or tesla units, data format, service request action, EOI action, and reset system to defaults.

Two analog outputs are available. These are a buffered version of the raw probe signal, and are not corrected for linearity or temperature.

All models are available without display or front buttons, for true “Black box” requirements.

Panel mount versions are available, and panels with correct cut-outs and mounting hardware.

The operating temperature range of the probe and instrument is 0 to 50 (C

DTM-150 Specifications

DTM-150	         performance summary      when combined with the listed probe.

| probe | active area (mm) | resolution of display | max field | finest resolution | accuracy at 25(C | tempco ppm/(C | zero drift (T/(C |
| --- | --- | --- | --- | --- | --- | --- | --- |
| LPT-130 | 4 x 1.6 | 1 in 600,000 | 3T | 1(T | (0.03% | -80 | (12 |
| LPT-230 | 4 x 1.6 | 1 in 600,000 | 0.3T | 0.1(T | (0.03% | -620 | (8 |
| MPT-132 | 1 x 0.5 | 1 in 600,000 | 3T | 1(T | (0.03% | -140 | (40 |
| MPT-230 | 1 x 0.5 | 1 in 600,000 | 0.3T | 0.1(T | (0.03% | -620 | (8 |

The following specifications refer to  	DTM-150 with LPT-130 probe   (standard size probe)

DTM-150 with MPT-132 probe  (miniature probe)

Ranges	4 ranges, selectable from front panel

0.3,   0.6,   1.2,   3.0 tesla full scale

Resolution	depends on the range selected, and whether the digital filtering is enabled.

Resolution available at the serial/GPIB port is greater than that of the display.

Resolution  ((T)

| Range selected | Filtering | display | serial/GPIB |
| --- | --- | --- | --- |
| 0.3 tesla | filtering ON | 1 | 0.1 |
|  | OFF | 5 | 0.1 |
| 0.6 tesla | filtering ON | 2 | 1 |
|  | OFF | 10 | 1 |
| 1.2 tesla | filtering ON | 4 | 1 |
|  | OFF | 20 | 1 |
| 3 tesla | filtering ON | 10 | 1 |
|  | OFF | 50 | 1 |

Absolute

Accuracy	(0.03% of reading   ( 0.006% of range     at 25(C

Temperature	with LPT-130 probe

Stability		scale factor:	-80 ppm of reading / (C  max.

zero drift:	( (12(T + 0.0003% of full scale ) / (C    max.

effect of probe cable: add -3ppm / (C for each metre of probe cable

with MPT-132 probe

scale factor:	-140 ppm of reading / (C  max.  (100ppm/ (C typical)

zero drift:	( (40(T + 0.0003% of full scale ) / (C    max.

effect of probe cable: add -3ppm / (C for each metre of probe cable

Frequency 	DC, or in AC mode  8 to 3,000 Hz

Update rate	10 fully corrected measurements per second

Display	7 digit, red LED.      8 additional indicators for range, units, peak hold and filter

Enclosure	All metal,   with tilt stand on bench models, bezel mount on rack models

Dimensions	217 x 125 x 50mm ,       weight 1.2 kg